# ANNEX II + III:	 TECHNICAL SPECIFICATIONS + TECHNICAL OFFER

# Contract title: Supply of  Medical Equipment >	p 1 /…

Publication reference: <24643/T3-K >

Columns 1-2 should be completed by the contracting authority

Columns 3-4 should be completed by the tenderer

Column 5 is reserved for the evaluation committee

Annex III - the contractor's technical offer

The tenderers are requested to complete the template on the next pages:

- Column 2 is completed by the contracting authority shows the required specifications (not to be modified by the tenderer),

- Column 3 is to be filled in by the tenderer and must detail what is offered (for example the words ‘compliant’ or ‘yes’ are not sufficient)

- Column 4 allows the tenderer to make comments on its proposed supply and to make eventual references to the documentation

The eventual documentation supplied should clearly indicate (highlight, mark) the models offered and the options included, if any, so that the evaluators can see the exact configuration. Offers that do not permit to identify precisely the models and the specifications may be rejected by the evaluation committee.

The offer must be clear enough to allow the evaluators to make an easy comparison between the requested specifications and the offered specifications.

| 1. Item number | 2. Specifications required | 3. Specifications offered | 4. Notes, remarks, ref to documentation | 5. Evaluation committee’s notes |
| --- | --- | --- | --- | --- |
|  | Operating table including positioning accessories -1 piece (Equipment for operating theatre) - Pillar - Electro-hydraulic height adjustment, backrest adjustment, Trendelenburg and Anti-Trendelenburg position, lateral tilt; - Height adjustment in the minimum range 680-1050mm; - Trendelenburg and Anti-Trendelenburg position adjustment in the minimum range -40 degrees/+40 degrees; - Lateral tilt adjustment in the minimum range -30 degrees/+30 degrees; - Remote control for table functions; - Table position memorization; - Built-in battery power supply; - Maximum load capacity minimum 350kg; - Power supply 220V, 50Hz; - Tabletop - Consists of minimum 5 segments; - Consisting of segments: head support, back section, pelvic section and divided leg section; - Translucent, for use with mobile X-ray; - Tabletop length 2000-2060 mm; - Tabletop width 500 - 550 mm; - Mattress resistant to disinfectants, waterproof, antistatic; - Mattress thickness minimum 80mm; 2.1 Back section - Tilt adjustment in the minimum range -40 degrees to +85 degrees; - Pelvic section consisting of one part 2.3 Head support - Tilt adjustment in the minimum range -50 degrees to +55 degrees; 2.4 Leg section - Consisting of two parts; - Tilt adjustment in the minimum range -90 degrees to +30 degrees; - Adjustments with an accuracy of 0,5mm and 0,1 degrees. - Accessories - IV Pole - Anaesthesia screen - High minimum 800 mm; - length of crossbar minimum 580mm; - Made of electropolished stainless steel; - Height adjustment and rotation rotation around its axis. - Anaesthesia tubes holder - to support anaesthetic tubes during surgery; - Adjustment of the height and distance of the support from the edge of the tabletop; - Support size 200 x 100 mm. - Armboard - to support the hand of a patient lying in a supine, prone or lateral position; - Dimensions minimum 160 x 500 mm; - rotate about its own axis by 180°; - Mounting straps for hand stabilisation. - Lateral support – 2 pcs - to support the patient's body in the lateral position; - The mounting system ensures stability even during extreme tilt when placing the patient in the lateral position; - Adjustment: height, distance from patient; - Dimensions minimum 152 x 70 x 24 mm. - Fixation belt - to secure the patient on the operating table in prone, supine and lateral positions; - Size adjustment by means of a hook&loop system; - Dimensions: strap length minimum 750 mm, width minimum 100 mm. - Leg holders, Goepel type - to support the patient's limb in lying or sitting position; - dimensions minimum 190 x 295 mm; - adjustment height, angle, rotation around its axis. - Gynaecological and urological bowl - to hold the body fluids during gynaecological or urological procedures; - with a drain; - dimensions minimum 65 x 260 x 320 mm; - capacity minimum 4 litres; - adjustment height, inclination angle, rotation around own axis. |  |  |  |
|  | Operating light with two light heads 1 piece Combination of two lights as one console for the operating room To have the possibility of electronic focusing. - It has control over the intensity of the brightness controlled through the display and through the sterile handle. - Adjusting the temperature to the color is controlled by the display and the sterile handle. - Endo mode with a green light controlled by the display and the sterile handle. - Setting on the Depth of illumination through the display and through the sterile handle. - Made of aluminum housing. - Integrated laser pointer. - System for automatic shadow management. - System for additional mechanical connection to the light field. - Control panel for setting the following minimum characteristics: illumination depth, shadow management, laser function activation/deactivation, endo illumination, illumination field size, illumination intensity, illumination color. - Control handles that have the option of sterilization, one for setting the focus on lighting, and the other with the option of setting for some of the functions: adjustment to brightness, adjustment to color to brightness, activation/deactivation to depth to light, activation/deactivation to laser function, activation/deactivation to endo illumination. 2 Main Light - Light intensity of at least 160,000 lux at 1 meter. - Diameter of the illuminated field d10 minimum 180 mm. - Diameter of the illuminated field d50 minimum 100 mm. - Option to set the brightness to 50-100%. - Ra-Ra color rendering index of at least 98. - Rendering index of color R9 R9 minimum 99. - Rendering index of R13 color minimum 99. - Possibility to set the minimum color temperature value in kelvin: 3750K; 4000K; 4250K; 4500K; 4750K. - Focusing light field of 18-34 cm. - Minimum of 5 Dimming Levels in the range of 50-100%. - Residual lighting with one shadow Min. 50%. - Residual lighting with two shadows Min. 55%. - Residual lighting on the bottom of the standardized tube minimum 100%. - Residual lighting on the bottom of a standardized tube and one shade Min. 50%. - Residual lighting on the bottom of a standardized tube and two shades Min. 55%. - Illumination depth 20 % Min. 2000 mm. - Illumination depth of 60% Min. 1050 mm. - Increase in temperature in the area of the head of a maximum of 0.5°C. - Working distance of 70-160cm. - Degree of protection of the light head: minimum IP-54. - Lamp life of at least 60,000 working hours. - Diameter of the light head: minimum 55 cm. - Maximum central radiation at a distance of 1 meter: no more than 560 W/m². 3 Satelite light - Light intensity of at least 160,000 lux at 1 meter. - Diameter of the illuminated field d10 minimum 180 mm. - Diameter of the illuminated field d50 minimum 100 mm. - Option to set the brightness to 50-100%. - Ra-Ra color rendering index of at least 98. - Rendering index of color R9 R9 minimum 99. - Rendering index of R13 color minimum 99. - Possibility to set the minimum color temperature value in kelvin: 3750K; 4000K; 4250K; 4500K; 4750K. - Focusing light field of 18-34 cm. - Minimum of 5 Dimming Levels in the range of 50-100%. - Residual lighting with one shadow Min. 50%. - Residual lighting with two shadows Min. 55%. - Residual lighting on the bottom of the standardized tube minimum 100%. - Residual lighting on the bottom of a standardized tube and one shade Min. 50%. - Residual lighting on the bottom of a standardized tube and two shades Min. 55%. - Illumination depth 20 % Min. 2000 mm. - Illumination depth of 60% Min. 1050 mm. - Increase in temperature in the area of the head of a maximum of 0.5°C. - Working distance of 70-160cm. - Degree of protection of the light head: minimum IP-54. - Lamp life of at least 60,000 working hours. - Diameter of the light head: minimum 55 cm. - Maximum central radiation at a distance of 1 meter: no more than 560 W/m². |  |  |  |
|  | Anaesthesia machine with patient monitor 1 piece - Able to work with all categories of patients (adults, children and neonates). - Mobile on wheels with a central brake. - Integrated aspirator with secretion jars. - Anesthetic gas and nitrous oxide evacuation system. - The gas evacuation system must be available as active and passive for the possibility of working with and without a compatible anesthetic gas evacuation system. - The active gas evacuation system must be connected to the anesthetic gas system with a control valve or ejector. - The passive gas evacuation system must be connected to a removal system with little or no suction flow with an overpressure valve and an underpressure valve. The maximum inlet flow must not exceed 0.5L/min. - Work with 3 gases: O2, AIR and N2O; with the possibility of mounting 2 spare cylinders at the back of the device. - The device must have an electronic gas meter. - Minimum 15´´ touch-sensitive screen that can simultaneously display 3 ventilation curves and numerical values ​​or optionally display 4 curves. - Integrated internal battery with a capacity of 120 minutes. - Breathing (circuit) system with an integrated heater, which can be disassembled without the need for tools. - Automatic self test (without user interaction) with calibration of all sensors and valves. - Operate in the following modes: conventional, low flow, and minimal flow anesthesia. - Soda lime absorber with a disposable CLIC technology canister with a capacity of max. up to 1.2 liters, as well as the possibility of a reusable canister with a capacity of 1.5 liters. - Able to perform mechanical ventilation using ambient air in the event of a shortage of all medical gases - to have an electric ventilator. - Independent of the central gas supply. - Automatic adjustment of alarm limits. - Pause function, for temporary interruption of ventilation and supply of fresh gas. - Integrated lighting on the work surface, with adjustment of the light intensity. - One drawer for storing material and which can be locked. - Command for rapid oxygen supply with a flow of 25-75 L/min. - The display shows the status of the device, with a display of the pressure in the airways, the status of the battery and the supply of gases from the central system and from the spare cylinders. - Ventilation modes: Manual; Spontaneous; Volume – controlled; Pressure – controlled; SIMV – VC; SIMV – PC; Autoflow, or compatible – delivery of a set volume at the minimum possible pressure; CPAP with pressure support. - Ventilator settings: Tidal volume 5-1500 ml with advanced neonatal support; Respiratory rate 3-100 bpm; Pinsp: from PEEP + 5 mbar to 80 mbar; Psupp: from 3 mbar to 80 mbar – PEEP; Inspiratory time: 0.2 - 10 sec; I:E ratio (inspiratory:expiratory) 1:50 to 50:1; Inspiratory flow minimum 0.1 to ≥ 160 LPM; Trigger sensitivity 0.3 - 15 l/min; Plateau time to inspiratory time 0 - 60%. - The device shall be supplied with a sevoflurane vaporizer, with the possibility of connecting a desflurane vaporizer. - Monitoring of the anesthesia device: Inspiratory O2 concentration with an O2 sensor with a 2-year lifespan and with monitoring of the lifespan of the oxygen cell; Tidal and minute volume; PIP, Pmean, Pplat, PEEP; Resistance (R) and Compliance (Cdyn); Optional simultaneous display of 2 Volume-Pressure and Flow-Volume loops. - Monitoring of anesthetic gases: Inspiratory and expiratory concentration of gases: O2, N2O, CO2 and anesthetics using a gas module, with display on the device display or on the hemodynamic monitor. - The device shall be supplied with at least 1 set of reusable adult silicone patient hoses that can be sterilized, anesthesia masks with sizes 3,4,5 that can be sterilized. Hemodynamic patient monitor supplied with the anesthesia machine: Screen with a minimum diagonal of 15 inches, touch-sensitive, with the ability to display 13 curves simultaneously; The patient monitor must be able to work with all categories of patients (adults, children and neonates); 150 hours of trends, 1200 NIBP measurements, 200 alarm events, 96h full disclosure; The patient monitor must have an integrated internal battery with a capacity of 5 hours; The monitor must have a centrally placed alarm signal lamp. Parameters that can be measured: NIBP; EKG 3 or 5 leads; 3 IBP; CO – cardiac output; temperature; SpO2; etCO2 (mainstream I sidestream); BIS (EEG curve, BIS trend, values: BIS, SQI, SR). Integrated three-channel printer. The following consumables are supplied with the monitor: One set of measurements: EKG (five leads), resp and HR; saturation (SpO2) with reusable adult sensor; NIBP for 3 reusable adult cuffs of different sizes; temperature with reusable sensor; CO2 sensor – capnography; invasive blood pressure measurement. |  |  |  |
|  | - Electrosurgical unit/device including accessories 1 piece - Touch screen with a minimum diagonal of 8 inches. - Automatic recognition of connected instruments. - Possibility of working in open surgery, laparoscopy and endoscopic surgery. - Minimum 6 individual sockets of which: Minimum two monopolar sockets; Minimum one bipolar socket; Minimum two universal sockets with the possibility of automatic recognition of the connected instrument; Minimum one socket for connecting a neutral electrode. - High frequency output power of minimum 500 kHz. - Contact Monitoring system for the neutral electrode with a signal for: unintentional removal of the neutral electrode by the patient, disconnection of the neutral electrode, short circuit of the neutral electrode. - To have a minimum of Bipolar modes: Bipolar cutting - Fast cutting with a small coagulation zone with a minimum power of 100W and a choice of at least 3 effects; Bipolar cutting for transurethral resection with a minimum power of 320W and a choice of at least 3 effects. - Cutting of various tissue structures, especially fibroid tissue with advanced bipolar technology with a minimum power of 200W and a choice of at least 3 effects. - Bipolar cutting with alternating output for fast and controlled cutting and enhanced smoke reduction with a minimum power of 300W and a choice of at least 3 effects. - Bipolar clean cutting for smooth and reproducible cutting in various tissues, e.g. muscle and fat with a minimum power of 200W and a choice of at least 3 effects. - Bipolar soft cutting with a minimum power of 200W and a choice of at least 3 effects. - Automatic coagulation - Slow bipolar coagulation for deep tissue effects, with auto stop function with a minimum power of 120W and a choice of at least 3 effects. - Automatic coagulation - Slow bipolar coagulation for deep tissue effects, with auto start function with a minimum power of 50W and a choice of at least 3 effects. - Strong bipolar coagulation with auto stop with a minimum power of 120W and a choice of at least 3 effects. - Coagulation of tissue with low adhesion and carbonization using advanced bipolar technology with a minimum power of 120W and a choice of at least 3 effects. - Bipolar soft coagulation with a minimum power of 120W and a choice of at least 3 effects. - Bipolar coagulation for transurethral resection with a minimum power of 200W and a choice of at least 3 effects. - Maximum output force during bipolar cutting ≥ 320W. - Maximum force during bipolar coagulation ≥ 200W. - To have at least monopolar modes: Monopolar cutting with an increased coagulation zone with a minimum power of 200W; Monopolar fine cutting for fast cutting with a small coagulation zone with a minimum power of 120W. - Monopolar superficial coagulation with increased dissection capabilities with a minimum power of 120W and a choice of at least 3 effects. - Monopolar clean cutting with a minimum power of 300W and a choice of at least 3 effects. - Monopolar non-contact coagulation for fulguration and superficial tissue effect with a minimum power of 120W. - Monopolar forced coagulation with a minimum power of 120W and a choice of at least 3 effects. - Monopolar strong superficial coagulation with increased dissection capabilities with a minimum power of 120W and a choice of at least 3 effects. - Monopolar gentle coagulation with a minimum power of 200W and a choice of at least 5 effects. - Monopolar spray coagulation with a minimum power of 120W and a choice of at least 3 effects. - A mode that allows ligation of blood vessels up to 7mm. - Maximum output power during monopolar cutting ≥ 300W. - Maximum power during monopolar coagulation ≥ 200W. - Possibility of upgrading with ultrasonic energy. - Appropriate trolley on the wheels for offered device. - Foot control with two pedals for activation (cutting/coagulation). - Set of at least 100 return electrodes, cable for connecting to the device. |  |  |  |
|  | Syringe pump 1 piece - Syringe infusion pump designed for target-controlled infusion. - At least 4,3-inch color screen pump with 16 different colors. - At least 96 combinations of colors and patterns. - Visual installation instructions on display. - Syringe visible during infusion + Syringe guard. - Flow rate setting range: - 0,01 - 150 mL/h (for 5 mL syringes). - 0,01 - 300 mL/h (for 10, 20 and 30 mL syringes). - 0,01 - 1200 mL/h (for 50/60 mL syringes). - Flow rate increment: - 0,01 mL/h step (0.01 to 10.00 mL/h). - 0,10 mL/h step (10.00 to 100.00 mL/h). - 1,00 mL/h step (100.00 to 1200.00 mL/h). - Compatible disposables Luer lock syringes, sizes: 5, 10, 20, 30, 50/60 mL - Dose mode: mL/h, μg/kg/min, mg/kg/h, Drug Library. - At least 10 Occlusion levels (±10 to ±120kPa). - KVO- Configurable: 0.1 mL/h - 3.0 mL/h, 0.1 mL/h step. - Alarms: Occlusion, Pressure, Nearly Empty, Slider Displacement, Syringe Barrel Detection, Syringe Displacement, Plunger Displacement, Battery, Re-alarm, Start Reminder, No Flow Rate, No VTBI, Flow Rate/VTBI Volume Judgment, Completion, Link Interruption, Power failure, Shutdown Notice. - Power supply AC 100-240V, 50-60 Hz. - Internal battery (Lithium-ion battery), Sub-battery (NiMH). - Main battery operating time Approx. 12 hours (at 5 mL/h and 25°C with a new fully charged battery). - TCI mode- 3 different views (quick display change): graph screen, text screen, -information screen according to healthcare professional needs. - Different plasma and effect-site models: - Propofol: Schnider, Marsh, Kataria, Paedfusor, - Remifentanil: Minto, - Sufentanil: Gepts, - Plasma limit setting in effect-site TCI. - Limitations of extreme parameters. - Flow rate limit settings based on patient characteristics. - Automatic and customizable graph axis scales. - Integrated handle. - Flexible pole clamp. - Possibility of linking devices using pumps station. - At least IP 24 protection level. |  |  |  |
|  | Surgical instruments Abdominal Surgical Instruments; Quantity. Scalpel handle No. 4, length 135mm; 4. Scalpel handle No. 4, length 214mm; 4. Surgical scissors, length 145mm, straight, sharp/blunt; 2. Surgical scissors, length 145mm, curved, sharp/blunt; 2. Surgical scissors, fine, length 150mm, Jameson, straight; 4. Surgical scissors, fine, length 180mm, Jameson, straight; 4. Tungstencut surgical scissors, length 145mm, straight, sharp/blunt; 2. Tungstencut surgical scissors, length 145mm, straight, sharp; 2. Surgical dissecting scissors, Lexer, length 165mm, curved, blunt; 2. Surgical dissecting scissors, Wertheim, length 200mm, curved, blunt; 4. Surgical dissecting scissors, Wertheim, length 200mm, curved, blunt; 4. Textile holder, Backhaus, length 115mm; 24. Textile holder, , length 265mm, Maier, with locking mechanism; 12. Dissection forceps, length 200mm, straight and curved; 12. Tissue forceps, length 200mm, straight, 1x2 teeth; 8. DeBakey vascular forceps, length 240mm, straight, width 2.0mm; 8. Pean surgical instrument, length 145mm, straight; 16. Pean surgical instrument, length 205mm, straight; 24. Hemostatic instrument, length 205mm, straight, 1x2 teeth; 20. Hemostatic instrument, length 200mm, curved, 1x2 teeth; 8. Vein retractor, length 200mm, 10x14mm; 4. Set of three Roux retractors, double-sided, length 165mm,; 4. Volkman retractor, with 6 sharp hooks, length 220mm,; 12. Kocher concave retractor, length 250mm, 80x65mm; 8. Farabeuf retractor, length 150mm, set of two retractors with dimensions 23x16mm, 26x16mm; (27x16mm, 30x16mm); 8. Carbide tipped needle holder, according to Mayo-Hegar, length 240mm, grip 0.5mm; 12. Schroeder grasping instrument, length 250mm, straight, 2x2; 8. Spatula, length 300mm, 37/45mm; 4. Babcock surgical tissue grasping instrument, length 175mm; 12. Collin surgical tissue grasping instrument, length 200mm, width 14mm; 8. Balfour abdominal retractor set, with extension 180mm, lateral blades 100x35mm, central blade 100x65mm; 4. Metal suction cannulas, length 240mm, diameter 10mm; 4. Metal liquid containers, 250ml, 95x51mm; 6. Metal mesh basket for washing instruments, 480x250x100mm; 2. Metal sterilization container with biological valve, 580x280x205mm; 2. Urology Surgical Instruments; Quantity. Randall Forceps 1/4 Curved 9 1/4"; 1. Randall Forceps 1/2 Curved 8 1/2"; 1. Randall Forceps 3/4 Curved 7 3/4"; 1. Randall Forceps Full Curved 7 3/4"; 1. Lees Bronchus Clamp 9"; 1. Van Buren Sound Set Of 17; 1. Otis Bougie A Boule 8fr; 1. Otis Bougie A Boule 12fr; 1. Otis Bougie A Boule 16fr; 1. Otis Bougie A Boule 10fr; 1. Wertheim-Cullen Pedicle Clamp 8 1/2"; 2. Ballenger Sponge Forceps Straight Serrated 7"; 2. Allis Tissue Forceps Delicate 5x6 9 1/2"; 2. Allis Tissue Forceps Delicate 5x6 7 1/2"; 2. Allis Tissue Forceps 5x6 6"; 2. Metzenbaum-Nelson Scissors Straight 9"; 2. Metzenbaum-Nelson Scissors Curved 9"; 2. Operating Scissors 5 1/2" Straight Sharp/Blunt; 2. Operating Scissors 6" Straight Sharp/Blunt; 2. Operating Scissors 6 1/2" Straight Sharp/Blunt; 2. Mayo-Hegar Needle Holder Serrated 10 1/2; 2. Gillies Skin Hook 7" 4mm; 1. Crile-Wood Needle Holder Narrow Jaws 6"; 1. Mayo-Hegar Needle Holder Serrated 6"; 1. Mayo-Hegar Needle Holder Serrated 7"; 1. Mayo-Hegar Needle Holder Serrated 8"; 1. Heaney Needle Holder Curved Serrated 8 1/4" TC; 1. Rochester Ochsner Forceps Curved 1x2 6 1/4; 1. Rochester Ochsner Forceps Straight 1x2 6 1/4"; 1. Rochester Ochsner Forceps Curved 1x2 7 1/4; 1. Rochester Ochsner Forceps Straight 1x2 7 1/4; 1. Roux Double-Ended Retractors Set/3; 1. Baby Farabeuf Retractors Set of 2 4 3/4"; 1. Backhaus Towel Clamp 5 1/4"; 4. Johns Hopkins Bulldog Clamp 2 1/2; 2. Tissue Forceps 1x2 5 1/2"; 2. Dressing Forceps Serrated 7"; 1. Tissue Forceps 1x2 7"; 1. Tissue Forceps 1x2 8"; 1. Tissue Forceps 1x2 10"; 1. Tissue Forceps 1x2 12"; 1. Tissue Forceps 3x4 5 1/2"; 1. Tissue Forceps 4x5 5 1/2"; 1. Tissue Forceps 3x4 6"; 1. Tissue Forceps 1x2 5 1/2"; 1. Tissue Forceps 1x2 6"; 1. Babcock Forceps 6 1/4"; 1. Babcock Forceps 8"; 1. Babcock Forceps 9 1/2"; 1. Rochester Pean Forceps Straight 6 1/4"; 1. Mixter Forceps Right Angled 7 1/2"; 1. Gemini Clamp Delicate Angled 8"; 1. Halsted Mosquito Forceps Straight 5"; 1. Halsted Mosquito Forceps Curved 5"; 1. Crile Forceps Straight 5 1/2"; 1. Crile Forceps Straight Delicate 5 1/2"; 1. Crile Forceps Curved Delicate 5 1/2"; 1. Rochester Pean Forceps Delicate Curved 8"; 1. Rochester Pean Forceps Straight 8"; 1. Mixter Forceps Curved 9 1/4"; 1. Rochester Pean Forceps Curved 9"; 1. Stille-list Bone Forceps Straight 10 1/2"; 1. Stille-list Bone Forceps Curved 10 1/2"; 1. Scalpel Handle #4; 2. Scalpel Handle #3; 2. Metal mesh basket for washing instruments, 480x250x100mm; 2. Metal sterilization container with biological valve, 580x280x205mm; 2. |  |  |  |
|  | Surgical suction pump 1 piece - Surgical aspirator with electric drive, with operating voltage 230V. - Suction power of 45l/min. - To enable continuous long-term work without interruption for more than 12 hours. - To be placed on a stand with 4 wheels, 2 of which with brakes. - To contain additional protection against the penetration of liquids into the device. - Maximum noise level of 48dB. - To be regulated and controlled by a foot pedal. - Integrated pressure gauge for precise display of the vacuum, as well as a button for its unlimited regulation. - Vacuum measurement 1......0 bar. - Maximum vacuum up to -91 kPa. - To have two 3-liter secretion containers, with an appropriate lid. - To contain 10 bacteriological filters with integrated overflow protection. - To contain suction hoses with appropriate adapter and hose holder. - All parts to be able to be sterilized in an autoclave at 134 degrees. - Power consumption does not exceed 175 W. |  |  |  |
|  | ICU beds -4 pieces - Electric hospital bed. - All plastic and painted elements are made with the addition of silver nanoparticles. - Dedicated panel for a three phases early mobilization program. - Side rails according to an international standard (EN 60601-2-52). - One-button activated positions: anti-shock, examination, Fowler, CPR, vascular, cardiac chair position. - Locking of individual function of the hand remote control or patient controls using the central panel. - Easy access to the CPR level. - Removable headboard. - X-ray translucent HPL bed surface. - Bed length extension. - Unlocked wheels alarm. - Low battery warning. - Length 2150-2200 mm (+ 320-330 mm bed length extension). - Outer width (with side rails) maximum 990 mm. - Mattress thickness minimum 160 mm in a waterproof, steam permeable cover. - Height regulation (min - max) 390 - 800 mm - Maximum backrest angle 70 degrees. - Maximum thigh segment angle 40 degrees. - Trendelenburg tilt minimum 16 degrees - Reverse-Trendelenburg tilt minimum 18 degrees. - Backrest autoregression 120 mm. - Autoregression of the thigh section 45 mm - Ingress protection (according to IEC 60529 standard) minimum IPX4, control and power supply system protected in IPX6 standard. - Maximum load minumum 280 kg. - Wheel diameter 150 mm. - Accessories: Drip bottle holder; Urinal holder; Patient chart holder, A4 paper format; Lifting pole with height adjustable holder. |  |  |  |
|  | Syringe pump-4 pieces - Syringe infusion pump designed for continuous or intermittent application of drugs. - At least 4,3-inch color screen pump with 16 different colours. - At least 96 combinations of colours and patterns. - Visual installation instructions on display. - Flow rate setting range: - 0,01- 150 mL/h (for 5 mL syringes), - 0,01- 300 mL/h (for 10, 20 and 30 mL syringes), - 0,01- 1200 mL/h (for 50/60 mL syringes). - Flow rate increment: - 0,01 mL/h step (0.01 to 10.00 mL/h), - 0,10 mL/h step (10.00 to 100.00 mL/h), - 1,00 mL/h step (100.00 to 1200.00 mL/h), - Compatible disposables Luer lock syringes, sizes: 5, 10, 20, 30, 50/60 mL. - Dose mode: mL/h, μg/kg/min, mg/kg/h, Drug Library. - Advanced dose mode: Interval/intermittent mode, multi-step mode, Delayed start mode. - KVO- Configurable: 0.1 mL/h - 3.0 mL/h, 0.1 mL/h step. - Drug library (up to 3000 drug entries, up to 30 user defined profiles) - Alarms: Occlusion, Pressure, Nearly Empty, Slider Displacement, Syringe Barrel Detection, Syringe Displacement, Plunger Displacement, Battery, Re-alarm, Start Reminder, No Flow Rate, No VTBI, Flow Rate/VTBI, Link Interruption, Power failure, Shutdown Notice. - Power supply AC 100-240V, 50-60 Hz - Internal battery: Lithium-ion battery. - Main battery operating time Approx. 12 hours (at 5 mL/h and 25°C with a new fully charged battery). - Integrated handle. - Flexible pole clamp. - Possibility of linking devices using pumps station. - At least IP 24 protection level. |  |  |  |
|  | Infusion pump-4 pieces - Volumetric infusion pump designed for continuous or intermittent application of drugs. - At least 4,3-inch color screen pump with 16 different colors. - At least 96 combinations of colors and patterns. - Visual installation instructions on display. - Flow rate: - When not using the drip sensor 0,10 to 1200,00 mL/h, - When using the drip sensor - 0,10 to 1200,00 mL/h (when setting is 20 drops/mL) - 0,10 to 300,00 mL/h (when setting is 60 drops/mL) - Flow rate increment: - 0,10 mL/h step (0,10 to 100,00 mL/h), - 1,00 mL/h step (100,00 to 1200,00 mL/h). - 10 occlusion levels (±30 to ±120kPa). - Dose mode: mL/h, μg/kg/min, mg/kg/h, Drug Library. - Advanced dose mode: Interval/intermittent mode, multi-step mode, Delayed start mode. - Combination dose mode: Solo and Pair Piggyback mode. - KVO- Configurable: - 0,10 to 3,00 mL/h (when not using the drip sensor (optional accessory)) - When using the drip sensor: - When using an infusion set of 20 drop/mL: 3,00 mL/h, - When using an infusion set of 60 drop/mL: 1,00 to 3,00 mL/h. - Drug library: up to 3000 drug entries, up to 30 user defined profiles. - Alarms: Upper Occlusion alarm, Lower Occlusion alarm, Nearly Empty alarm, Pressure alarm, Air-in-line alarm, Door alarm, Flow Rate Abnormality alarm, Free Flow alarm, Line Empty alarm, Drip Sensor Dislocation alarm, Battery alarm, Shutdown Notice alarm, Power Failure alarm, Re-alarm, Start Reminder, No Flow Rate alarm, No VTBI alarm, Completion alarm, Link Interruption alarm. - Alarms display with realistic illustrations. - Anti free flow function. - Power supply AC 100-240V, 50-60 Hz. - Internal battery. Main battery operating time Approx. 5 hours (at 25 mL/h and 25°C with a new fully charged battery). - Integrated handle. - Flexible pole clamp. - Possibility of linking devices using pumps station. - At least IP 22 protection level. - Use of peristaltic method that allows partial occlusion of the infusion tubing, to ensure more precise and stable fluid delivery. - Includes a mechanism that ensures minimal flow rate fluctuations during prolonged infusions, especially at low flow rates. |  |  |  |
|  | Patient monitor -4 pieces - LCD or LED color monitor with flat screen, touch sensitive. - Minimum size 13 inches diagonal. - Minimum resolution 1920x1080 pixels. - Monitoring range: adult, pediatric and neonatal. - Must have at least 10 waveforms. - For monitoring at least basic parameters: 3/5/12 channel ECG, HR, SpO2, RESP, NIBP, PR, 2xTEMP. - Minimum calculation of medications, ventilation, oxygenation, hemodynamics, renal function. - Analysis of at least 30 types of arrhythmias. - Alarms for all parameters (technical and physiological). Minimum three levels of alarms with different light and sound intonation. - Individuals setting limits for each alarm. - Night mode of operation. - Definition of colors of curves displayed by the user. - Have a minimum of 3 types of filters: monitor, diagnostic and surgical. - Power supply 220-230V / 50 – 60Hz. - Battery power supply minimum 5100 mAh to allow a minimum of 8 hours of operation. - ECG monitoring selection: 3 leads I, II, III, 5 leads I, II, III, aVR, aVF, aVL V, and 10 leads I, II, III, aVR, aVF, aVL, V1-V6 with selection of tracking speed in the minimum range 6.25-50 mm/s. - ECG signal acquisition with a minimum frequency of 1000 Hz. - ST segment analysis. - QT/QTc measurements. - Measure heart rate HR: in the minimum range 15-300 bpm for adults and 15-350 bpm for children/neonates, accuracy of ±1%. - Pulse measurement (via saturation) from 20-300 bpm, accuracy of ±2 bpm. Possibility of measuring pulse also via NIBP signal. - Body temperature measurement from 0-50 ˚С with accuracy of ±0.2 ˚С. - NIBP measurement: diastolic minimum range 10-250 mmHg, systolic minimum range 25-290 mmHg, with minimum resolution of 1 mmHg. - RESP measurement in minimum range 0-200 rpm with accuracy of ±2 rpm. - Internal memory for storing at least the following parameters: 240 hours of trend review (graph/table), 1600 blood pressure measurements. - Minimum 2xUSB connectors for connecting to other equipment or transferring data from internal memory. - RJ-45 connector for connecting to a network with a central monitoring system. - Degree of protection against liquid penetration minimum IPX1. - Built-in protection against the use of a defibrillator. - Possibility of connecting a monitor to a central monitoring station. - In accordance with at least the following IEC 60601-1; IEC 60601-1-2; EN 60601-1; EN 60601-1-2; IEC 80601-2-49. - Suitable trolley on wheels for the monitor. |  |  |  |
|  | Mechanical ventilation device - Intensive care ventilator - Has to be designed to work with invasive, non-invasive ventilation and oxygen therapy. - Color screen, with a minimum size of 12 inches, touch sensitive. - Be able to ventilate patient categories: infants, pediatric and adults. - The device itself should have its own cart, with lockable wheels. - Provide compressed air drive through its own built-in turbine that operates independently of the central compressed air system. - It should allow spontaneous breathing without stress with excellent activation response time thanks to its own turbine. - To be used for inter-clinical patient transport and in an ambulance without the addition of an additional compressor. - Emergency valve in the event of an oxygen supply interruption, automatically switching the device to spontaneous breathing mode through filtered ambient air. - The battery should be internal with a built-in independent power supply and a minimum operation time of 45 min. as well as optionally having the possibility of upgrading with an additional battery for an additional minimum of 5 hours of operation. - To have an output for connecting a pneumatic medication nebulizer that is synchronized with inspiration. - Software procedure for easy and fast aspiration, whereby by pressing a button, the device switches to suction mode, and after completion, an automatic valve is turned on that oxygenates the patient for at least 2 minutes. - The respirator must have the following ventilation modes: - Volume controlled ventilation modes: VC-CMV / VC-AC, VC-SIMV, VC-MMV; pressure-controlled ventilation modes: PC-APRV, PC-BIPAP / PC-SIMV+; PC-AC. - Spontaneous breathing support: SPN-CPAP. - To have O2 therapy using masks for non-invasive ventilation. - Non-invasive ventilation (NIV) available in all modes with a very fast response time to the patient's efforts and automatic leak compensation. - Advanced option of volume-oriented ventilation modes that deliver the set volume, under the lowest possible pressure or equivalent (PRVC, AutoFlow, ASV). - To allow ventilation of babies starting from 5kg body weight. - Simultaneous display on the screen of 3 curves (pressure, volume, flow) trends. - The respirator should have the following characteristics: Inspiratory pressure 1 – 99 mbar, Ventilation frequency 2 – 80 per minute, O2 concentration adjustment in the range 21 – 100%, Inspiratory hold function for X-ray imaging, Trigger sensitivity adjustment 1 – 15 L/min, Support for spontaneous breathing of the patient (Plessure support) 0 – 50 mbar, Maximum inspiratory flow 250 L/min, Inspiration time 0,2 – 10 seconds (duration of the inspiratory phase), Respiratory (tidal) volume 20ml – 2L, PEEP from 0 – 50 mbar. - Leak and trigger compensation, as well as automatic tube compensation. - Low pressure oxygen (LPO) inlet for ventilation, without central gas supply. - Parameters to be displayed: Pressure (maximum, plateau, average), Total MV, MV for spontaneous breathing, Tidal volume VT, inspiratory and expiratory VT, Respiratory rate (total and spontaneous), FiO2, Inspiratory gas temperature. - Loop display: Pressure-Volume, Volume-Flow, Flow-Pressure, Volume-CO2, Ptrach-Volume, Flow-Ptrach. - To have separate functions for inspiratory hold and expiratory hold (Insp. Hold; Exp. Hold). - PC232 communication interface. - Nurse call port. - Measurement and calculation of the patient weaning factor from mechanical ventilation: RSBI. - To have measurements of: R - resistance, C - compliance, Intrinsic PEEP/autoPEEP. - Use of standard patient hoses and to have alarms by priority. - Materials to be delivered with the device: Holder joint for attaching the patient hose, Nebulizer for medication, 50 pcs. of disposable patient hoses for adults with a length of 1.5-1.8 meters. |  |  |  |